## Slide 1: Short Title: GC-MS Uncertainty PI: J. Tyler Davidson Sam Houston State University Assessing the uncertainty of GC-MS chromatographic peak area: Collaboration with IAB Members (CBP and FBI)
Student Participants: Emma K. Hardwick

## Slide 2: IAB Member Involvement
FBI - Marc LeBeau
CBP - Jonathan McGrath, Patricia Hawes, Tom Schoch
LLNL - Audrey Williams
Teams Meetings:
  March 12th, May 29th, June 24th
Email Correspondence:
  May 31st (CBP requests Tom Schoch), June 24th (proposal and summary of progress to date sent to Audrey Williams and Tom Schoch)
Input:
  Inclusion of retention time uncertainty
  Addition of metabolites (4-ANPP & 6-MAM)
  Removal of splitless component and only using two concentrations
  Incorporation of LLNL in update meetings
  Consider the effect of different integrators as a sub-study

## Slide 3: IAB Research Need
Understanding the Uncertainty Associated with the Peak Area of a GC-MS Chromatographic Peak
Gas chromatography-mass spectrometry (GC-MS)
  Widely used analytical technique (qualitative/quantitative)
Forensic science
  Impact of incorrect analytical determinations
  Critical to understand uncertainty – assessment and communication are difficult
GC-MS uncertainty of measurement
  Retention time, ion abundance, and peak area
Goal - assess the sources of uncertainty and their corresponding magnitudes on GC-MS chromatographic peak area

## Slide 4: Updated Reference Mixture
Range of molecular weights and polarities
Required to assess uncertainty based on physiochemical properties
Initial proposed composition based on the 2022 NFLIS-Drug annual report
Modified to meet specific interests of the IAB members
Concentrations:
  Low – PCP (12.5 ppm), fentanyl (25 ppm), MDMA, pentylone, cocaine, 4-ANPP, xylazine (50 ppm), everything else 250 ppm
  High – PCP (12.5 ppm), fentanyl (50 ppm), everything else 500 ppm
Analytical reference mixture composition |  | 
Early | Middle | Late
Amphetamine | Xylazine | 6-MAM
Methamphetamine | Cocaine | Heroin
MDMA | Codeine | Fentanyl
Pentylone | 4-ANPP | Clonazepam
PCP | ∆9-THC | Bromazolam

## Slide 5: Updated Experimental Design
Assessed variables
  Injection type – 10:1 split vs. splitless (15:1)
  Concentration – low vs. high
  Elution time – early-, middle-, late-elution
  Instrument usage – based on the number of injections
  Instrument – 2021, 2014, 2009
Completed GC-MS maintenance overhaul to establish baseline conditions
  Cleaned EI source
  Filaments
  DB-5MS column
  Gold seal
  Inlet liners
  Septa
Routine maintenance schedule for commonly replaced consumables
  Septum – weekly
  Inlet liner + O-ring – monthly
  Syringe – every 3 months

## Slide 6: Progress Update
Optimized chromatographic separation of target analytes (19.5 mins)
Amphetamine, methamphetamine, and MDMA basic extraction
Standardized ChemStation integrator parameters for automated integration
Calculated expanded uncertainty of measurement
  2σ – 2 x %RSD of replicate measurements (approximates a 95% CI)
  Within-day, within-week, within-month
5 replicates per day x 2 concentrations x 3 instruments
Completed 8 days, 5 weeks, 1 month of data collection

## Slide 7: Initial Expanded Uncertainty Results
Within-day (n=8)
Spooled = pooled 2*%RSD (i.e., ~95%CI)
Variables:
  Instrument
  Concentration
  Elution
  Injections (TBD)
PA – Instrument, Concentration, Elution(?)
RPA > PA
RRT ≈ RT
Repeated for w/in week and w/in month

## Slide 8: Expanded Uncertainty Tool

## Slide 9: Challenges Encountered
Budget issues
  Funds have not been received
  Resolution - progress to date has been completed using internal funds
Technical issues
  Slightly behind projected timeline due to financial delays and impacts to university access (i.e., Hurricane Beryl)
  Resolution – plenty of time to collect sufficient data to address the research question
Will there be any protection of IP related to this effort? If so, describe the IP
  Not applicable

## Slide 10: Planned Publications/Presentations
American Academy of Forensic Sciences (AAFS)
  CARFS reception in 2025
  Conference presentation in 2026
Peer-reviewed manuscript

## Slide 11: Questions?
IAB member questions?
Our questions for IAB members
  Feedback on the expanded uncertainty tool?
  Additional information beyond agency, data, and analyst?